## FV1 2025
Decommissioning Equipment Form
*This list is supplemental material to the Core Biorisk Management Documents provided by Sandia National Laboratories. It provides a general task list for equipment decommissioning. If any elements below are not required or not applicable, document as such or remove from form as needed.
Device Name: |  |  |  | ID Number:
Final disposition of equipment: | ÿ Storage: ÿ Trade-in: ÿ Donate: ÿ Discard/Recycle
| Location | Manufacturer |  | Receiver
Date of Disposition:
Task | Responsibility | Due Date | Completion Date | Comments
GENERAL PREPARATION
Determine if equipment will be relocated, stored, shipped, donated, or discarded.
Document the decommissioning process for auditing and regulatory compliance.
Device tested for functionality prior to storage. |  |  |  | ÿ Functional ÿ NOT Functional, explain:
Determine most effective decontamination method to inactivate any of the pathogens ever used in the equipment. Note: may require including qualified outside contractors.
Computer data downloaded and stored. No protected health information (PHI) or other identifiable data remains.
GENERAL SAFETY
Wear appropriate PPE for any biological, chemical, radiological, or other hazards that may be encountered during equipment decommissioning. |  |  |  | List PPE:
Evaluate electrical safety (unplug, internal power systems, etc.).
Handle sharp parts with appropiate precautions to prevent punctures/lacerations. |  |  |  | Sharps mitigation:
Decontaminate device prior to removing from a containment area. (See Decontamination section below.)
Disconnect any gas/air lines or pressurized sources.
Additional hazards (radiological, laser, lithium power source, hazardous metals, etc.). Check device manual for information. |  |  |  | Describe additional hazards:
Task | Responsibility | Due Date | Completion Date | Comments
DECONTAMINATION
Follow manufacturer's instructions for decontamination (if available). |  |  |  | ÿ Not available ÿ Source:
Prepare disinfectant solution that is compatible with device and effective for the type of hazards in use and in accordance with manufacturer's instructions. |  |  |  | ÿ Disinfectant used and contact time:
Remove all consumables and dispose as directed by facility waste disposal policies and procedures, including sharps disposal.
Drain any liquids from the equipment prior to moving.
Clear device of any debris.
Place removable parts into disinfection solution for the appropriate contact time. Label and log any removable parts for accountability. Do not submerge electrical components or other items if prohibited by the manufacturer.
Disinfect the exterior of the device, including any trays, tubing, lines, etc.
For cold storage (freezer/refrigerator): ÿ Empty and defrost. ÿ Disinfect exterior of tube/vial/box and triple package for transport if relocating. ÿ Inactivate any samples as needed by submerging in appropriate solution (for disposal). ÿ Disinfect stored items such as cold packs, freezer boxes, containers, and dispose of as appropriate.
Decontaminate device using a compatible method:
ÿ steam sterilization ÿ ethylene oxide, or ÿ gas plasma ÿ vaporous hydrogen peroxide ÿ low temp/dry-heat method: ____________
Indicate decontamination validation process and documentation. |  |  |  | ÿ PASS ÿ FAIL Indicator type:
Record decontamination results and document for regulatory compliance.
Biohazard markings can be removed after proper decontamination.
Task | Responsibility | Due Date | Completion Date | Comments
DISPOSITION: Storage
Device and parts have been decontaminated, and are clean and dry.
All device parts are present and accounted for.
Device is carefully boxed/bagged with all parts, user manual, and associated maintenance records (or location of records are indicated). |  |  |  | Location of documents:
Exterior of box/bag is labeled with device name, serial number, date of decommissioning and owner/department.
Device is stored in a secure location (for dual-use item)1 or in a location where it can be monitored (closet, storage room, etc.). |  |  |  | Storage location:
Establish a schedule and log to check and document device condition on a regular basis. |  |  |  | Schedule of checks:
DISPOSITION: Trade-in to Manufacturer
Contact manufacturer representative for trade-in options. |  |  |  | Contact information:
Obtain Return Material Authorization or approval for shipment of current device. |  |  |  | Authorization number: Date received:
Confirm device and parts are decontaminated, clean, and dry.
Account for all device parts and attachments.
Confirm that device is functional. |  |  |  | ÿ Functional ÿ NOT Functional, explain:
Obtain shipping materials to secure and protect the device.
Remove any labels, hazard identification, property tags (return according to organization's policies and procedures). |  |  |  | ÿ Not applicable ÿ Property tag return date:
Complete documentation required by the manufacturer for the return.
Package and ship the device according to manufacturer's instructions. |  |  |  | Carrier & Shipping ID:
Retain copies of all documentation for reference. |  |  |  | Location of documents:
DISPOSITION: Donate
The condition, functionality and capabilities of the device are clearly communicated to the receipient. |  |  |  | Recipient info:
Consult with BRM Advisor to determine if device is a dual-use item. If so, suggest avoiding donation.
Device and parts are decontaminated, clean, and dry.
Account for all device parts and attachments.
Confirm that device is functional. |  |  |  | ÿ Functional ÿ NOT Functional, explain:
Compile the device's history, documentation, and all available information on its condition to send to recipient. Do not disclose sensitive information related to use of high-consequence pathogens or other biosecurity concerns. |  |  |  | ÿ User Manual ÿ Maintenance logs ÿ Service contact/warrantee
Copy and store the device's history, documentation, maintenance records, etc. according to organizational requirements. |  |  |  | Stored as: ÿ Hard-copy ÿ Electronic / scanned files ÿ Other : Storage Location:
Secure device for shipping or for safe transport if within facility. |  |  |  | Shipping/transport details:
Task | Responsibility | Due Date | Completion Date | Comments
DISPOSITION: Discard/Recycle
Obtain authorization to discard equipment, consulting local, regional, and national regulations and well as intstiutional policy for disposal. |  |  |  | Authorized by: Date:
Notify BRM Advisor of intent to discard if dual-use equipment. |  |  |  | ÿ Not applicable ÿ Contacted, date:
If dual-use item, functionality has been permenantly disabled. |  |  |  | ÿ Not applicable ÿ Describe how disabled:
Account for all device parts and attachments.
Remove any labels, hazard identification, property tags (return according to organization's policies and procedures). |  |  |  | ÿ Not applicable ÿ Property tag return date:
Store the device's history, documentation, maintenance records, etc. according to organizational requirements. |  |  |  | Stored as: ÿ Hard-copy ÿ Electronic / scanned files ÿ Other : Storage Location:
Arrange for disposal/recycling according to organization's policies and procedures. |  |  |  | Date of disposal: Additional details:
AUTHORIZATIONS
SUPERVISOR
(Department name) | Printed name | Date |  | Signature
BRM ADVISOR
| Printed name | Date |  | Signature
FACILITIES MANAGER
| Printed name | Date |  | Signature
1Dual-use items: Certain materials, information and technologies that are intended for benefit, but which might be misapplied to do harm. (Laboratory Biosafety Manual 4th edition, WHO, 2020)